## Slide 1: The New High Accuracy Default PTP Profile in the IEEE 1588 Draft Revision
Maciej Lipiński
European Organization for Nuclear Research

## Slide 2: Agenda
Overview
High Accuracy Default PTP Profile
  Optional features
  Inter-operation with other Default PTP Profiles
  Model of Local PTP Clock
Summary
2

## Slide 3: Agenda
Overview
High Accuracy Default PTP Profile
  Optional features
  Inter-operation with other Default PTP Profiles
  Model of Local PTP Clock
Summary
2

## Slide 4: White Rabbit
Update of CERN’s control and timing system
Sub-ns accuracy & sub-10ps precision
Based on IEEE 802.3, 802.1Q and 1588 (PTP) with extensions:
  L1 synchronization, similar to SyncE
  Precise phase detection
  Calibration and online estimation of asymmetries
  Extension of the PTP protocol to accommodate the above
De-facto standard for timing & daq in scientific applications
Implementation
3

## Slide 5: IEEE1588 standard revision
4
2013: start of IEEE1588 revision
Project Authorisation request: “…support for synchronization to better than 1 nanosecond”
5 sub-committees
High Accuracy sub-committee
  Focus on White Rabbit
  Experts from industry and academia
  Division of WR into self-contained parts
  Definition of Optional Features and PTP Profile that allow WR-like implementation and WR performance

## Slide 6: Agenda
Some history
High Accuracy Default PTP Profile
  Optional features
  Inter-operation with other Default PTP Profiles
  Model of Local PTP Clock
Summary

## Slide 7: High Accuracy Default PTP Profile
High Accuracy Delay Request-Response Default PTP Profile
5

## Slide 8: High Accuracy Default PTP Profile
High Accuracy Delay Request-Response Default PTP Profile
5

## Slide 9: High Accuracy Default PTP Profile
High Accuracy Delay Request-Response Default PTP Profile
Identification: 00-1B-19-01-01-00
PTP Attributes: default values & ranges
Optional features /mechanism
  BMCA: default
  Mechanism: request-response by default, peer-to-peer allowed
  Required options: none
  Prohibited options: none
  Permitted options: all*
Clock physical requirements: frequency accuracy below 0.01%
* Optional feature inactive (disabled) by default
5

## Slide 10: High Accuracy Default PTP Profile
Identification: 00-1B-19-02-01-00
PTP Attributes: default values & ranges
Optional features /mechanism
  BMCA: default
  Mechanism: request-response by default, peer-to-peer allowed
  Required options: Annex O, Q.3, 16.7, 16.8, 17.6*, masterOnly*
  Prohibited options: 16.5
  Permitted options: all others*
Clock physical requirements: frequency accuracy below 4.6ppm
Default values & ranges for attributes of required options
High Accuracy model of Local PTP Clock
Information on inter-operation with other Default PTP Profiles
* Optional feature inactive (disabled) by default
High Accuracy Delay Request-Response Default PTP Profile
5

## Slide 11: Agenda
Some history
High Accuracy Default PTP Profile
  Optional features
  Inter-operation with other Default PTP Profiles
  Model of Local PTP Clock
Summary

## Slide 12: Options required/prohibited by HA Profile
  Required and active by default:
  L1-based sync. performance enhancements [Annex O]
  Configurable correction of timestamps [16.7]
  Calculation of the <delayAsymmetry> for certain media [16.8]
  Requirements of Q.3 [Annex Q]
  Required and inactive by default:
  Mechanism for external config. of a PTP Instance’s PTP Port state [17.6]
  masterOnly mode [9.2.2.2 & 8.2.15.5.2]
  Prohibited:
  Isolation of PTP Instances [16.5]
6

## Slide 13: Options required/prohibited by HA Profile
  Required and active by default:
  L1-based sync. performance enhancements [Annex O]
  Configurable correction of timestamps [16.7]
  Calculation of the <delayAsymmetry> for certain media [16.8]
  Requirements of Q.3 [Annex Q]
  Required and inactive by default:
  Mechanism for external config. of a PTP Instance’s PTP Port state [17.6]
  masterOnly mode [9.2.2.2 & 8.2.15.5.2]
  Prohibited:
  Isolation of PTP Instances [16.5]
6

## Slide 14: L1 synchronization enhancements (L1Sync)
  Supports cooperation & defines relationship between
    PTP-based synchronization and
    L1-based syntonization, e.g. Synchronous Ethernet
  Provides configuration and status of the PTP-L1 relationship
  IMPORTANT: Requires hardware support
7

## Slide 15: PTP synchronization and L1 syntonization
PTP synchronization
clkLocalPTP_A=clkLocalPTP_B
L1 syntonization
clkL1x_A=clkL1x_B
L1 syntonization
clkL1x_A=clkL1x_B
PTP Port/Instance A
PTP Port/Instance B
.
Local PTP
Clock
Local PTP
Clock
8

## Slide 16: PTP synchronization and L1 syntonization
PTP synchronization
clkLocalPTP_A=clkLocalPTP_B
L1 Syntonization
clkL1x_A=clkL1x_B
L1 Syntonization
clkL1x_A=clkL1x_B
PTP Port/Instance A
PTP Port/Instance B
  Transmit coherent port: xtx = constant, so clkL1tx= clkLocalPTP
  Receive coherent port : xrx = constant, so clkL1rx= clkLocalPTP
  Congruent port : flow of L1 syntonization and PTP synchronization are the same
Local PTP
Clock
Local PTP
Clock
8

## Slide 17: PTP synchronization and L1 syntonization
PTP synchronization
clkLocalPTP_A=clkLocalPTP_B
L1 Syntonization
clkL1x_A=clkL1x_B
L1 Syntonization
clkL1x_A=clkL1x_B
PTP Port/Instance A
PTP Port/Instance B
  Transmit coherent port: xtx = constant, so clkL1tx= clkLocalPTP
  Receive coherent port : xrx = constant, so clkL1rx= clkLocalPTP
  Congruent port : flow of L1 syntonization and PTP synchronization are the same
Local PTP
Clock
Local PTP
Clock
8

## Slide 18: PTP synchronization and L1 syntonization
PTP synchronization
clkLocalPTP_A=clkLocalPTP_B
L1 Syntonization
clkL1x_A=clkL1x_B
L1 Syntonization
clkL1x_A=clkL1x_B
PTP Port/Instance A
PTP Port/Instance B
  Transmit coherent port: xtx = constant, so clkL1tx= clkLocalPTP
  Receive coherent port : xrx = constant, so clkL1rx= clkLocalPTP
  Congruent port : flow of L1 syntonization and PTP synchronization are the same
Local PTP
Clock
Local PTP
Clock
8

## Slide 19: PTP synchronization and L1 syntonization
PTP synchronization
clkLocalPTP_A=clkLocalPTP_B
L1 Syntonization
clkL1x_A=clkL1x_B
L1 Syntonization
clkL1x_A=clkL1x_B
PTP Port/Instance A
PTP Port/Instance B
  Transmit coherent port: xtx = constant, so clkL1tx= clkLocalPTP
  Receive coherent port : xrx = constant, so clkL1rx= clkLocalPTP
  Congruent port : flow of L1 syntonization and PTP synchronization are the same
Local PTP
Clock
Local PTP
Clock
8

## Slide 20: L1 synchronization enhancements (L1Sync)
  Data sets: storing config & status of the L1-PTP relationship (congruent, coherent)
  L1_SYNC TLVs:
    Detection of L1Sync-supporting PTP Ports
    Exchange of config/status/parameters
  FSM: control of establishing of the required PTP-L1 relationship
  Know PTP-L1 relationship: allows correction of timestamp with phase measurement
9

## Slide 21: L1Sync in High Accuracy Default PTP Profile
PTP Port/Instance A
PTP Port/Instance B
  Requirement for L1Sync PTP Port to be: - transmit and receive coherent
  - congruent
xtx_A = xrx_B = xtx_B = known constant
xrx_A = phase measured to enhance timestamp precision
Local PTP
Clock
Local PTP
Clock
10

## Slide 22: Options required/prohibited by HA Profile
  Required and active by default:
  L1-based sync. performance enhancements [Annex O]
  Configurable correction of timestamps [16.7]
  Calculation of the <delayAsymmetry> for certain media [16.8]
  Requirements of Q.3 [Annex Q]
  Required and inactive by default:
  Mechanism for external config. of a PTP Instance’s PTP Port state [17.6]
  masterOnly mode [9.2.2.2 & 8.2.15.5.2]
  Prohibited:
  Isolation of PTP Instances [16.5]

## Slide 23: Configurable correction of timestamps
11
  IEEE1588-2019 (and 2008) defines timestamping ref. plane and allows correction of timestamps for <ingressLatency> and <egressLatency>:
  <ingressTimestamp>=<ingressProvidedTimestamp> – <ingressLatency>
  <egressTimestamp> = <egressProvidedTimestamp> + <egressLatency>
Ref.
plane
PTP
code
PTP message
<ingressLatency>
<egressLatency>
PTP device
Medium

## Slide 24: Configurable correction of timestamps
11
  IEEE1588-2019 (and 2008) defines timestamping ref. plane and allows correction of timestamps for <ingressLatency> and <egressLatency>:
  <ingressTimestamp>=<ingressProvidedTimestamp> – <ingressLatency>
  <egressTimestamp> = <egressProvidedTimestamp> + <egressLatency>
  This optional feature:
    Defines timestampCorrectionPortDS dataset to store the values of <egressLatency> & <egressLatency>
    Mandates correction of timestamps with the provided values

## Slide 25: Options required/prohibited by HA Profile
  Required and active by default:
  L1-based sync. performance enhancements [Annex O]
  Configurable correction of timestamps [16.7]
  Calculation of the <delayAsymmetry> for certain media [16.8]
  Requirements of Q.3 [Annex Q]
  Required and inactive by default:
  Mechanism for external config. of a PTP Instance’s PTP Port state [17.6]
  masterOnly mode [9.2.2.2 & 8.2.15.5.2]
  Prohibited:
  Isolation of PTP Instances [16.5]

## Slide 26: Calculation of <delayAsymmetry>
  IEEE1588-2019 (and 2008) provides calculation of <meanDelay> and allows correction of <meanDelay> for <delayAsymmetry>: <meanDelay> = [(t2 – t1 ) + (t4 – t3 )] / 2 tms = <meanDelay> + <delayAsymmetry>
  tsm = <meanDelay> – <delayAsymmetry>
  This optional feature:
    Defines <delayCoefficient> (α) to characterize asymmetry of applicable bidirectional media (e.g. bidirectional fiber)
    Specifies <delayAsymmetry> calculation using <delayCoefficient> (α)
    <delayAsymmetry> = constantAsymmetry + [α/(α+2)]·(<meanDelay>)
    Defines asymmetryCorrectionPortDS dataset for <delayCoefficient> & constantAsymmetry
    Mandates correction of <meanDelay> with the calcualted <delayAsymmetry>
12

## Slide 27: Options required/prohibited by HA Profile
  Required and active by default:
  L1-based sync. performance enhancements [Annex O]
  Configurable correction of timestamps [16.7]
  Calculation of the <delayAsymmetry> for certain media [16.8]
  Requirements of Q.3 [Annex Q]
  Required and inactive by default:
  Mechanism for external config. of a PTP Instance’s PTP Port state [17.6]
  masterOnly mode [9.2.2.2 & 8.2.15.5.2]
  Prohibited:
  Isolation of PTP Instances [16.5]

## Slide 28: Calibration
Values of <ingressLatency>, <egressLatency>, <delayCoefficient> obtained using Calibration Procedures [Annex Q]
[Q.3] defines requirements for PTP Instance in calibration procedures:
  Output signal, e.g. pulse-per-second (PPS)
  Optional features in 16.7 & 16.8 supported and enabled
  The value <meanDelay> exposed to the user
  Timestamps with precision sufficient for the intended accuracy of synchronization
  <egressProvidedTimestamp> and <ingressProvidedTimestamp> corrected for semi-static latency (e.g. bitslide, see Q.2 and P.4.2)
  Timestamping Clock is the Local PTP Clock
13

## Slide 29: Options required/prohibited by HA Profile
  Required and active by default:
  L1-based sync. performance enhancements [Annex O]
  Configurable correction of timestamps [16.7]
  Calculation of the <delayAsymmetry> for certain media [16.8]
  Requirements of Q.3 [Annex Q]
  Required and inactive by default:
  Mechanism for external config. of a PTP Instance’s PTP Port state [17.6]
  masterOnly mode [9.2.2.2 & 8.2.15.5.2]
  Prohibited:
  Isolation of PTP Instances [16.5]

## Slide 30: External config. of PTP Port state & masterOnly
Best Master Clock Algorithm (BMCA) defines the state of each PTP Port in BC/OC and synchronization hierarchy in the entire PTP Network
Mechanism for external config. of a PTP Instance’s PTP Port state:
  Disables BMCA and affects all PTP Ports on a PTP Instance
  Allows configuring the desired state of each PTP Port
  Defines externalPortConfigurationPortDS dataset to store the desired state
  Provides external configuration of synchronization hierarchy in the entire PTP network
masterOnly mode on PTP Port:
  Maintains operation of BMCA and provides configuration per PTP Port
  Results in PTP Port ignoring Announce messages and defaulting to Master state
  Prevents PTP Instance from “slaving” on particular PTP Ports (e.g. access ports)
14

## Slide 31: Options required/prohibited by HA Profile
  Required and active by default:
  L1-based sync. performance enhancements [Annex O]
  Configurable correction of timestamps [16.7]
  Calculation of the <delayAsymmetry> for certain media [16.8]
  Requirements of Q.3 [Annex Q]
  Required and inactive by default:
  Mechanism for external config. of a PTP Instance’s PTP Port state [17.6]
  masterOnly mode [9.2.2.2 & 8.2.15.5.2]
  Prohibited:
  Isolation of PTP Instances [16.5]

## Slide 32: Agenda
Some history
High Accuracy Default PTP Profile
  Optional features
  Inter-operation with other Default PTP Profiles
  Model of Local PTP Clock
Summary

## Slide 33: Inter-operation with Default PTP Profiles
High Accuracy Delay Request-Response Default PTP Profile
  Inter-operation with Delay Request-Response Default PTP Profile
  Inter-operation with Peer-to-Peer Default PTP Profile, if peer-to-peer mechanism supported
  No inter-operation with Transparent Clocks
  No enhancements of synchronization performance
High Accuracy Default PTP profile
Request-Response
Default PTP Profile
Operation as if it was Request-Response Default PTP Profile
High Accuracy Default PTP profile
PTP Port in Master state:
Request-Response
Default PTP Profile
PTP Port in Slave state:
Operation as if connected to a Request-Response Default PTP Profile
15

## Slide 34: Agenda
Some history
High Accuracy Default PTP Profile
  Optional features
  Inter-operation with other Default PTP Profiles
  Model of Local PTP Clock
Summary

## Slide 35: High Accuracy model of Local PTP Clock
High Accuracy Local PTP Clock in J.5.5 – highlights:
Syntonization of Local PTP Clock
  Physical to the Grandmaster, e.g. L1 clock signal
  Independent from the PTP timing message exchange
Synchronization of Local PTP Clock using PTP timing messages
16

## Slide 36: High Accuracy model of Local PTP Clock
16

## Slide 37: Agenda
Some history
High Accuracy Default PTP Profile
  Optional features
  Inter-operation with other Default PTP Profiles
  Model of Local PTP Clock
Summary

## Slide 38: Summary
IEEE1588
“Core”
Defines attributes
and mechanisms
17
High Accuracy Request-Response Default PTP Profile
(Annex J)

## Slide 39: Summary
HA-specific optional features
(active by default)
IEEE1588
“Core”
Defines attributes
and mechanisms
Enhance
17
High Accuracy Request-Response Default PTP Profile
(Annex J)
Requires and specifies
default/allowed values
L1 Sync
Configurable correction of timestamps
Calculation of the delayAsymmetry

## Slide 40: Summary
HA-specific optional features
(active by default)
IEEE1588
“Core”
Defines attributes
and mechanisms
Enhance
17
High Accuracy Request-Response Default PTP Profile
(Annex J)
Requires and specifies
default/allowed values
L1 Sync
Configurable correction of timestamps
Calculation of the delayAsymmetry
Calibration
(Annex Q)

## Slide 41: Summary
Generic optional features
(inactive by default)
HA-specific optional features
(active by default)
IEEE1588
“Core”
Mechanism for external configuration
master Only
mode
Defines attributes
and mechanisms
Enhance
17
Modify
BMCA
Requires and specifies
default/allowed values
Requires and specifies
default/allowed values
High Accuracy Request-Response Default PTP Profile
(Annex J)
L1 Sync
Configurable correction of timestamps
Calculation of the delayAsymmetry
Calibration
(Annex Q)

## Slide 42: Summary
Generic optional features
(inactive by default)
HA-specific optional features
(active by default)
Sub-ns synchronization using High Accuracy Default Profile
(Annex P)
Informative annex describes HA Profile
implementation that provides
sub-ns accuracy of synchronization
IEEE1588
“Core”
Mechanism for external configuration
master Only
mode
Defines attributes
and mechanisms
Requires and specifies
default/allowed values
Requires and specifies
default/allowed values
Enhance
17
Modify
BMCA
High Accuracy Request-Response Default PTP Profile
(Annex J)
L1 Sync
Configurable correction of timestamps
Calculation of the delayAsymmetry
Calibration
(Annex Q)

## Slide 43: Summary
Generic optional features
(inactive by default)
HA-specific optional features
(active by default)
Sub-ns synchronization using High Accuracy Default Profile
(Annex P)
Informative annex describes HA Profile
implementation that provides
sub-ns accuracy of synchronization
IEEE1588
“Core”
Mechanism for external configuration
master Only
mode
Defines attributes
and mechanisms
Requires and specifies
default/allowed values
Requires and specifies
default/allowed values
Enhance
17
Modify
BMCA
High Accuracy Request-Response Default PTP Profile
(Annex J)
L1 Sync
Configurable correction of timestamps
Calculation of the delayAsymmetry
Calibration
(Annex Q)
???

## Slide 44: Backup slides

## Slide 45: PTP Profiles and Optional Features
18

## Slide 46: PTP Profiles and Optional Features
“Synchronization framework”
“Extensible &
parameterized
core”
(attributes, delay mechanism)
Optional parts
(optional features)
18

## Slide 47: PTP Profiles and Optional Features
“Synchronization framework”
“Extensible &
parameterized
core”
(attributes, delay mechanism)
Optional parts
(optional features)
Transport mappings
Informative Annexes
18

## Slide 48: PTP Profiles and Optional Features
“Synchronization framework”
“Extensible &
parameterized
core”
(attributes, delay mechanism)
Optional parts
(optional features)
Transport mappings
Informative Annexes
PTP
Applications
Telecom
Power
18

## Slide 49: PTP Profiles and Optional Features
“Synchronization framework”
“Extensible &
parameterized
core”
(attributes, delay mechanism)
Optional parts
(optional features)
Profile: delay mechanism, attribute values & options
Transport mappings
Informative Annexes
PTP
Applications
Telecom
Power
18

## Slide 50: PTP Profiles and Optional Features
“Synchronization framework”
“Extensible &
parameterized
core”
(attributes, delay mechanism)
Optional parts
(optional features)
Profile: delay mechanism, attribute values & options
Transport mappings
Informative Annexes
PTP
Applications
Telecom
Power
ITU-T Profiles
IEEE/IEC Profiles
18

## Slide 51: PTP Profiles and Optional Features
“Synchronization framework”
“Extensible &
parameterized
core”
(attributes, delay mechanism)
Optional parts
(optional features)
PTP
Applications
Default Profiles
Delay Request-Response
Peer-to-peer
High Accuracy Delay Request-Response
Profile: delay mechanism, attribute values & options
Transport mappings
Informative Annexes
18